## reg_array_segments
BCM Address map segments table by function
Address range is part of Struck SIS8300-L2 Address Map from address 0x400 to 0xFFF
ADDR
START | END | Segment Name
0x400 | 0x40F | BCM INTEGRATION
0x410 | 0x41F | BCM CORE CONTROL
0x420 | 0x42F | BCM CORE STATUS
0x430 | 0x43F | BCM CORE MPS CONTROL
(n=5=ch1, n=6=ch2, n=E=ch10)
0xn00 | 0xn0F | BCM ADC CONTROL CH(n-4)
0xn10 | 0xn1F | BCM ADC MEASUREMENTS CH(n-4)
0xn20 | 0xn2F | BCM DIFF CONTROL (n-4)
0xn30 | 0xn3F | BCM DIFF STATUS (n-4)
0xn40 | 0xn4F | BCM_FIBER_OUT_DATA_SELECTn (n-4)
0xn50 | 0xn5F | BCM_PROBE_SOURCE_SELECTn (n-4)
NOTE: "n" number of ADC, DIFF, FIBER OUT, PROBE channels is not necessarily the same and depends on the actual HW and Desy core configuration!
Current configuration for Struck is
#define NUM_ADC_CHAN 10
#define NUM_PROBE_CHAN 4
#define NUM_DIFF_CHAN 10
#define NUM_FIBER_IN_CHAN 2
#define NUM_FIBER_OUT_CHAN 2

## reg_array_regster_description
Addr | Acces | Name | Bit Description | Related to this registers in previous BCM version | Register Name
BCM INTEGRATION
0x400 | RO | BCM_INT_GATEWARE_VERSION | Gateware version: bit0-7 - Minor(8bit) bit8-15 - Major(8bit) bit16-31 - Device ID(16bit) | 0x400 | #define BCM_INT_GATEWARE_VERSION_REG 0x400 // RO
0x401 | RO | BCM_INT_SOFTWARE_VERSION | Has value 0x00000000 at powerup | 0x401 | #define BCM_INT_SOFTWARE_VERSION_REG 0x401 // RO
0x402 | RO | BCM_INT_MIN_TRIG_PERIOD_SRC_SEL | bit0: select source for the minimal trigger value and maximum beampulse width 0 - source is integration register (defalult) 1 - source is MRF timing interface |  | #define BCM_INT_MIN_TRIG_PERIOD_SRC_SEL_REG 0x402 // RO
0x404 | R/W | BCM_INT_TRIG_SYNC_DELAY | bit0-4: fixed number of delay taps for IODELAY module for delaying trigger signal |  | #define BCM_INT_TRIG_SYNC_DELAY_REG 0x404 // R/W
0x405 | R/W | BCM_INT_TRIGR_SYNC_DELAY_CTRL | IODELAY module control (not yet defined) |  | #define BCM_INT_TRIGR_SYNC_DELAY_CTRL_REG 0x405 // R/W
BCM CORE CONTROL
0x410 | R/W | BCM_CTRL_RFQ_TRANSPARENCY | bit0-15: percentage of beam arriving after the RFQ due to the RFQ beam loss |  | #define BCM_CTRL_RFQ_TRANSPARENCY_REG 0x410 // R/W
0x411 | AC | BCM_CTRL_DIFF_WARN_RESET | bit0: Reset for min/max values in FAST, MEDIUM FAST and SLOW transmission interlcoks |  | #define BCM_CTRL_DIFF_WARN_RESET_REG 0x411 // AC
0x412 | R/W | BCM_CTRL_MIN_TRIG_PERIOD | bit0-15 : Min. trigger period (1/frequency) - equals the min. beam pulse period for all ADC channels, range 71 ms to 65 seconds |  | #define BCM_CTRL_MIN_TRIG_PERIOD_REG 0x412 // R/W
0x413 | R/W | BCM_CTRL_MAX_BEAMPULSE_WIDTH | bit0-12: Max. beam pulse width, same limit for all channels. Beam pulse width is calculated from the number of samples above 3 mA. Also the trigger pulse width is checked against this limit, range 5 us to 8.191 ms |  | #define BCM_CTRL_MAX_BEAMPULSE_WIDTH_REG 0x413 // R/W
0x414 | R/W | BCM_CTRL_ENABLE_CALIBRATION_PULSE | bit0: calibration pulse enable 0 - pulse disabled 1 - pulse enabled |  | #define BCM_CTRL_ENABLE_CALIBRATION_PULSE_REG 0x414 // R/W
BCM CORE STATUS
0x420 | RO | BCM_STATUS | bit0 : BCM core busy status signal 0 – core not busy, input parameters can change 1 – BCM core busy, input parameters can not change, any write to BCM core registers will be ignored and bit1 of this register will be set bit1 : previous write status 0 – previous write to BCM register space was accepted, BCM core was not busy 1 - previous write to BCM register space was NOT accepted, BCM core WAS busy bit4: 1- main clock is missing or corrupted bit5: 1 - auxiliary clock is missing or corrupted |  | #define BCM_STATUS_REG 0x420 // RO
0x421 | RO | BCM_MEAS_CLK_FREQ_KHZ | Bit0-19 : Measured frequency of the main clock [kHz] |  | #define BCM_MEAS_CLK_FREQ_KHZ_REG 0x421 // RO
0x422 | RO | BCM_MEAS_TRIG_PERIOD | Bit0-31: Measured main trigger period (max. 45 sec) [1/88MHz] |  | #define BCM_MEAS_TRIG_PERIOD_REG 0x422 // RO
0x423 | RO | BCM_MEAS_TRIG_WID | Bit0-19: Measured main trigger pulse width (max. 11 ms) [1/88MHz] |  | #define BCM_MEAS_TRIG_WID_REG 0x423 // RO
0x424 | RO | BCM_MEAS_TIME_FLATTOP | Bit0-19: Duration of time window as defined for flattop calculation |  | #define BCM_MEAS_TIME_FLATTOP_REG 0x424 // RO
BCM CORE MPS CONTROL
0x430 | AC | BCM_MPS_RESET | bit0: Reset "latched" alarms (BCM_MPS_HOLD_x registers) and first alarm readback (BCM_MPS_FIRST_x registers) | 0x404 | #define BCM_MPS_RESET_REG 0x430 // AC
0x431 | R/W | BCM_MPS_ALL_ALARM_DISABLE | bit0: 0 - all alarms enabled, but can be individually disabled in BCM_CLK_TRIG_MPS_ALARM_DISABLE, BCM_ADC_MPS_ALARM_DISABLE and BCM_DIFF_MPS_ALARM_DISABLE registers 1 - disable all alarms on all channels | 0x403 | #define BCM_MPS_ALL_ALARM_DISABLE_REG 0x431 // R/W
0x432 | R/W | BCM_CLK_TRIG_MPS_ALARM_DISABLE | individual alarm disable for clock and trigger alarm bits bit0 : 0 - auxiliary clock (125Mhz) measurement alarm active 1 - alarm disabled bit1: 0 - main clock (88MHz) measurement alarm active 1 - alarm disabled bit2: 0 - alarm active for Trigger width shorter than announced by Timing System 1 - alarm disabled bit3: 0 - alarm active for Trigger width longer than 4.5us (4.5us is set in BCM PKG_TYPES.vhd) 1 - alarm disabled bit4: 0 - alarm active for check of the Trigger period less than limit from Timing System 1 - alarm disabled |  | #define BCM_CLK_TRIG_MPS_ALARM_DISABLE_REG 0x432 // R/W
0x433 | RO | BCM_CLK_TRIG_MPS_ALARM_HOLD | MPS "latched" alarms readback: bits which are '1' show alarms activated after the last MPS reset bit0 : 0 - Measurement of the auxiliary clock (125Mhz) based on main clock (88MHz) is IN range 1 - Measurement of the auxiliary clock (125Mhz) based on main clock (88MHz) is OUT of range bit1: 0 - Measurement of the main clock (88MHz) based on auxiliary clock (125MHz) is IN range 1 - Measurement of the main clock (88MHz) based on auxiliary clock (125MHz) is OUT of range bit2: 0 - Trigger width (5 us .. 6 ms) is ok 1 - Trigger width (5 us .. 6 ms) is longer than announced by Timing System bit3: 0 - Trigger width is ok 1 - Trigger width is less than 4.5us (4.5us is set in BCM PKG_TYPES.vhd) bit4: 0 - Trigger period (71 ms .. 30 s) is ok 1 - Trigger period (71 ms .. 30 s) is less than limit from Timing System | 0x421 | #define BCM_CLK_TRIG_MPS_ALARM_HOLD_REG 0x433 // RO
0x434 | RO | BCM_CLK_TRIG_MPS_ALARM_FIRST | first alarm readback: bits which are '1' show the first activated alarm(s); reset by writing 1 to the BCM_MPS_RESET register bit0: Measurement of the auxiliary clock (125Mhz) based on main clock (88MHz) is out of range bit1: Measurement of the main clock (88MHz) based on auxiliary clock (125MHz) is out of range bit2: Trigger width (5 us .. 6 ms) is longer than announced by Timing System bit3: Trigger width is less than 4.5us (4.5us is set in BCM PKG_TYPES.vhd) bit4: Trigger period (71 ms .. 30 s) is less than limit from Timing System | 0x421 | #define BCM_CLK_TRIG_MPS_ALARM_FIRST_REG 0x434 // RO
BCM ADC CONTROL CHn (n=5=ch1, n=6=ch2, n=E=ch10)
0xn00 | R/W | BCM_ADC_TRIG_FINE_DELAY | bit0-10: Trigger fine delay value to align the pulse edges with the trigger edges for each ADC individually. signed, 1/88M, range -11.6 .. +11.6 us | 0x406 | #define BCM_ADC_TRIG_FINE_DELAY_REG 0x00 // R/W
0xn01 | R/W | BCM_ADC_SCALE | bit0-15: Input scale factor to compensate amplification differences in the analogue input chains [unsigned, start value: approx. 25600] |  | #define BCM_ADC_SCALE_REG 0x01 // R/W
0xn02 | R/W | BCM_ADC_UPPER_THRESHOLD | bit0-15: Upper beam current threshold - active continuously. Negative values will result in continuous alarm; very high values will mask out the alarm completely; [signed, mA/256] |  | #define BCM_ADC_UPPER_THRESHOLD_REG 0x02 // R/W
0xn03 | R/W | BCM_ADC_LOWER_THRESHOLD | bit0-15: Lower beam current threshold - active during a defined interval based only on the trigger. With negative values, no alarm will be generated even without beam; very high values will always generate an alarm at every trigger. [signed, mA/256] |  | #define BCM_ADC_LOWER_THRESHOLD_REG 0x03 // R/W
0xn04 | R/W | BCM_ADC_ERRANT_THRESHOLD | bit0-15: Errant beam current threshold; with negative or close to zero values, an alarm will be generated even without beam; for very high values, never an alarm will be generated; [signed, mA/256] |  | #define BCM_ADC_ERRANT_THRESHOLD_REG 0x04 // R/W
0xn05 | R/W | BCM_ADC_OFFSET | bit0-15: ADC offset compensation; [signed, ADC_LSBs] |  | #define BCM_ADC_OFFSET_REG 0x05 // R/W
0xn06 | R/W | BCM_ADC_DROOPRATE | bit0-15 droop rate (depending on ACCT time constant), value of 32768 corresponds to a droop of 2.1%/ms [unsigned, 2^(-37)*88M] |  | #define BCM_ADC_DROOPRATE_REG 0x06 // R/W
0xn07 | R/W | BCM_ADC_ENABLE_PARAM_CHOPPER_RFQ | Configuration bits for ADCs, MEBT chopper and RFQ bit0: droop compensation and baseline correction enable bit, '0' - disabled, '1' - enabled) bit1: input noise filter (moving averager) enable bit; '0' - disabled, '1' - enabled bit4: beam pulse duration parameter; '0' - behind MEBT chopper, '1' - before MEBT chopper bit5: RFQ loss parameter; '0' - behind RFQ; '1' - before RFQ | 0x407, 0x408, 0x40A | #define BCM_ADC_ENABLE_PARAM_CHOPPER_RFQ_REG 0x07 // R/W
0xn08 | R/W | BCM_ADC_MPS_ALARM_DISABLE | alarm disable for ADC channels, 0 - alarm enabled, 1- alarm disabled bit0 - ADC CHn - upper current threshold bit1 - ADC CHn - lower current threshold bit2 - ADC CHn - errant beam threshold bit3 - ADC CHn - beam pulse longer than trigger bit4 - ADC CHn - beam pulse longer than limit bit5 - ADC CHn - overflow alarm inihibit for ADC channels bit6 - ADC CHn - underflow alarm inihibit for ADC channels bit7 - ADC CHn - adc stuck, value does not change | 0x407, 0x408, 0x40A | #define BCM_ADC_MPS_ALARM_DISABLE_REG 0x08 // R/W
BCM ADC MEASUREMENTS CHn (n=5=ch1, n=6=ch2, n=E=ch10)
0xn10 | RO | BCM_MEAS_ADC_OFFS_ERR_SHORT | bit0-15: ADC offset error averaged over the last 32 samples, [signed, ADC LSB] |  | #define BCM_MEAS_ADC_OFFS_ERR_SHORT_REG 0x10 // RO
0xn11 | RO | BCM_MEAS_ADC_OFFS_ERR_INTEG | bit0-23: ADC offset error integrated over a whole beam pulse period with beam, [signed, ADC_LSBs*2^19/88M] |  | #define BCM_MEAS_ADC_OFFS_ERR_INTEG_REG 0x11 // RO
0xn12 | RO | BCM_MEAS_TIME_TRIG_TO_PULSE | bit0-10: Measure time between rising trigger edge and rising beam pulse edg, [signed, 1/88M] |  | #define BCM_MEAS_TIME_TRIG_TO_PULSE_REG 0x12 // RO
0xn13 | RO | BCM_MEAS_DROOP_ERR | bit0-15: Measure error of droop compensation, [signed, @MW: unit tbd] |  | #define BCM_MEAS_DROOP_ERR_REG 0x13 // RO
0xn14 | RO | BCM_MEAS_BEAMPULSE_WIDTH | bit0-19: Measured beam pulse width; number of samples with amplitude above 3 mA, inside the "beam allowed" window, [unsigned, 1/88M] |  | #define BCM_MEAS_BEAMPULSE_WIDTH_REG 0x14 // RO
0xn15 | RO | BCM_MEAS_CHARGE | bit0-31: Charge of one pulse (accumulated samples inside the "beam allowed" window), [signed, mA/64/88M] |  | #define BCM_MEAS_CHARGE_REG 0x15 // RO
0xn16 | RO | BCM_MEAS_CHARGE_FLATTOP | bit0-31: Charge of one pulse (accumulated samples inside a specified window where the flattop is assumed), [signed, mA/64/88M] |  | #define BCM_MEAS_CHARGE_FLATTOP_REG 0x16 // RO
0xn17 | RO | BCM_MEAS_CALPULSE_AMPL_EARLY | bit0-15: Measured amplitudes of calibration pulse shortly after start of calibration pulse: starting 70 us after start of calibration pulse, [signed, mA/256] |  | #define BCM_MEAS_CALPULSE_AMPL_EARLY_REG 0x17 // RO
0xn18 | RO | BCM_MEAS_CALPULSE_AMPL_LATE | bit0-15: Measured amplitudes of calibration pulse shortly before end of calibration pulse, [signed, mA/256] |  | #define BCM_MEAS_CALPULSE_AMPL_LATE_REG 0x18 // RO
0xn19 | RO | BCM_ADC_MPS_ALARM_HOLD | latched alarm readback: bits which are '1' show alarms activated after the last MPS reset; reset by writing 1 to the BCM_MPS_RESET register bit0 - ADC CHn - upper current threshold bit1 - ADC CHn - lower current threshold bit2 - ADC CHn - errant beam threshold bit3 - ADC CHn - beam pulse longer than trigger bit4 - ADC CHn - beam pulse longer than limit bit5 - ADC CHn - overflow alarm inihibit for ADC channels bit6 - ADC CHn - underflow alarm inihibit for ADC channels bit7 - ADC CHn - adc stuck, value does not change |  | #define BCM_ADC_MPS_ALARM_HOLD_REG 0x19 // RO
0xn1A | RO | BCM_ADC_MPS_ALARM_FIRST | "upper current threshold" first alarm readback: bits which are '1' show the first activated alarm(s); reset by writing 1 to the BCM_MPS_RESET register bit0 - ADC CHn - upper current threshold bit1 - ADC CHn - lower current threshold bit2 - ADC CHn - errant beam threshold bit3 - ADC CHn - beam pulse longer than trigger bit4 - ADC CHn - beam pulse longer than limit bit5 - ADC CHn - overflow alarm inihibit for ADC channels bit6 - ADC CHn - underflow alarm inihibit for ADC channels bit7 - ADC CHn - adc stuck, value does not change |  | #define BCM_ADC_MPS_ALARM_FIRST_REG 0x1A // RO
BCM DIFF CONTROL n, (n=5=ch1, n=6=ch2, n=E=ch10)
0xn20 | R/W | BCM_DIFF_SELECT_CHAN_INPUT_A | bit0-3: select source for the differential channel input A. Available sources are ADC channels and fiber inputs |  | #define BCM_DIFF_SELECT_CHAN_INPUT_A_REG 0x20 // R/W
0xn21 | R/W | BCM_DIFF_SELECT_CHAN_INPUT_B | bit0-3: select source for the differential channel input B. Available sources are ADC channels and fiber inputs |  | #define BCM_DIFF_SELECT_CHAN_INPUT_B_REG 0x21 // R/W
0xn22 | R/W | BCM_DIFF_DELAY | bit0-10: delay to be added to one of the channels of a differential pair to compensate beam arrival time differences and different cable lengths; [signed, 1/88M], range -11.6 .. +11.6 us |  | #define BCM_DIFF_DELAY_REG 0x22 // R/W
0xn23 | R/W | BCM_DIFF_RELAX_AT_SLOPE | bit0-3: Relax parameter depending on the properties of the BCTs connected to the differential pair; [integer 0..15] @MW: more explanation necessary when implemented! |  | #define BCM_DIFF_RELAX_AT_SLOPE_REG 0x23 // R/W
0xn24 | R/W | BCM_DIFF_FAST_THRESH | bit0-14: Threshold for the FAST absolute value of beam current difference (weighted with the "Relax parameter") between the two channels of a differential pair, [unsigned, mA/256] | 0x405 | #define BCM_DIFF_FAST_THRESH_REG 0x24 // R/W
0xn25 | R/W | BCM_DIFF_MEDIUM_THRESH | bit0-31: Threshold for ACCUMULATED CHARGE DIFFERENCE over a moving time window defined by BCM_DIFF_MEDIUM_WIDTH1 and BCM_DIFF_MEDIUM_WIDTH2 registers; [unsigned, mA/32/88M] | 0x405 | #define BCM_DIFF_MEDIUM_THRESH_REG 0x25 // R/W
0xn26 | R/W | BCM_DIFF_MEDIUM_WIDTH_1 | bit0-4: Averaging window width factor 1 (as large as possible, should be > 15 for all cases when the window is > 15/88M); [integer] | 0x405 | #define BCM_DIFF_MEDIUM_WIDTH_1_REG 0x26 // R/W
0xn27 | R/W | BCM_DIFF_MEDIUM_WIDTH_2 | bit0-13: Averaging window width factor 2 (as small as possible); [integer] | 0x405 | #define BCM_DIFF_MEDIUM_WIDTH_2_REG 0x27 // R/W
0xn28 | R/W | BCM_DIFF_SLOW_THRESH | bit0-30: Threshold for ACCUMULATED CHARGE over the total time allowed by the "beam allowed" time window; [unsigned, mA/32/88M] | 0x405 | #define BCM_DIFF_SLOW_THRESH_REG 0x28 // R/W
0xn29 | R/W | BCM_DIFF_MPS_ALARM_DISABLE | alarm disable for DIFF channels, 0 - alarm enabled, 1- alarm disabled bit0 - DIFF CHn - Fast differential threshold bit1 - DIFF CHn - Medium differential threshold bit2 - DIFF CHn - Slow differential threshold | 0x405 | #define BCM_DIFF_MPS_ALARM_DISABLE_REG 0x29 // R/W
BCM DIFF STATUSn (n=5=ch1, n=6=ch2, n=E=ch10)
0xn30 | RO | BCM_DIFF_FAST_MIN | bit0-15: Minimum of the FAST difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/128] |  | #define BCM_DIFF_FAST_MIN_REG 0x30 // RO
0xn31 | RO | BCM_DIFF_FAST_MAX | bit0-15: Maximum of the FAST difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/128] |  | #define BCM_DIFF_FAST_MAX_REG 0x31 // RO
0xn32 | RO | BCM_DIFF_MEDIUM_MIN | bit0-15: Minimum of the MEDIUM FAST difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/16/88M] |  | #define BCM_DIFF_MEDIUM_MIN_REG 0x32 // RO
0xn33 | RO | BCM_DIFF_MEDIUM_MAX | bit0-15: Maximum of the MEDIUM FAST difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/16/88M] |  | #define BCM_DIFF_MEDIUM_MAX_REG 0x33 // RO
0xn34 | RO | BCM_DIFF_SLOW_MIN | bit0-15: Minimum of the SLOW difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/16/88M] |  | #define BCM_DIFF_SLOW_MIN_REG 0x34 // RO
0xn35 | RO | BCM_DIFF_SLOW_MAX | bit0-15: Maximum of the SLOW difference (as applied to the threshold comparator) since the last pulse of "warn_reset", enabled during "beam allowed" time window; [signed, mA/16/88M] |  | #define BCM_DIFF_SLOW_MAX_REG 0x35 // RO
0xn36 | RO | BCM_DIFF_MPS_ALARM_HOLD | latched alarm readback: bits which are '1' show alarms activated after the last MPS reset; reset by writing 1 to the BCM_MPS_RESET register bit0 - DIFF CHn - Fast differential threshold bit1 - DIFF CHn - Medium differential threshold bit2 - DIFF CHn - Slow differential threshold" |  | #define BCM_DIFF_MPS_ALARM_HOLD_REG 0x36 // RO
0xn37 | RO | BCM_DIFF_MPS_ALARM_FIRST | first alarm readback: bits which are '1' show the first activated alarm(s); reset by writing 1 to the BCM_MPS_RESET register bit0 - DIFF CHn - Fast differential threshold bit1 - DIFF CHn - Medium differential threshold bit2 - DIFF CHn - Slow differential threshold" |  | #define BCM_DIFF_MPS_ALARM_FIRST_REG 0x37 // RO
BCM FIBER OUT DATA SELECT (n=5=ch1, n=6=ch2, n=E=ch10)
0xn40 | R/W | BCM_FIBER_OUT_DATA_SELECT | bit0-5: select output data stream (88 MSPS) for fiber output ch, representing processed ADC data stream to be sent to another FPGA module for diff. comparison |  | #define BCM_FIBER_OUT_DATA_SELECT_REG 0x40 // RO
BCM PROBE SOURCE SELECT (n=5=ch1, n=6=ch2, n=E=ch10)
0xn50 | R/W | BCM_PROBE_SOURCE_SELECT | bit0-5: Select source for probe channel n (up to 64 possibilities per channel) to be stored in RAM |  | #define BCM_PROBE_SOURCE_SELECT_REG 0x50 // RO